## Slide 1
GEO-ISC meeting, Hannover
May 6th, 2011
  Towards a Suspension Platform Interferometer
  Katrin Dahl for the AEI 10 m Prototype team
AEI 10m Prototype

## Slide 2: Outline
A brief introduction to the AEI 10 m Prototype
Motivation Suspension Platform Interferometer
SPI
  Optical layout
  Construction
  Vacseal
  Pointing noise of fibre injectors
  Phase meter interface
Lessons learned

## Slide 3: AEI 10 m Prototype
The AEI 10 m Prototype will provide ultra-low displacement noise testing environment
  To probe at (and later go beyond) the standard quantum limit
  Prove new techniques, e.g. Khalili cavity
  For GRACE follow-on and LISA related experiments
  Train people for GEO-HF

## Slide 4: Vacuum system
Tubes:
1.5 m diameter
Volume ca. 100 m3
22 t stainless steel
Tanks:
3 m diameter, 3.4 m tall
Tank centers
separated by 11.65m
Roughing: One 170 l/s screw pump
Main pumps: Two 2000 l/s turbo-molecular pumps
Backing and differential pumping: Two scroll pumps
10-6 mbar after about 12 hours
10-7 mbar after about 2 weeks
of pumping

## Slide 5: Seismic Attenuation System
Modified version of HAM-SAS
AEI-SAS is a passive isolation system equipped with sensors and actuators for additional active stabilization
For all DOF resonance frequency below 0.4 Hz
Individual GAS filters 80 dB isolation, expected horizontal peak isolation 80 dB
March 4th, installation of the first table
However, some residual inter-table motion is left

## Slide 6: SPI & table actuation
The SPI is
a sensor to measure residual motions between the tables
These relative motions will then be stabilised by appropriate feedback to actuators
Actuators
voice-coil actuator at each GAS and at each inverted pendulum of the table
Vertical
actuator
Horizontal actuator
Suspension Platform Interferometer

## Slide 7: Why a Suspension Platform Interferometer?
Ease lock acquisition of cavities by reducing residual test mass motion
Reduction of burden to actuators on the suspended mirrors
Testbed for GRACE follow-on and LISA related experiments
   Sets requirements on SPI

## Slide 8: Optical layout of the SPI
Heterodyne Mach-Zehnder IFOs with unequal arm lengths (by 23m)
No specific lock point, enables to track table motion over more than one wavelength
Heterodyne frequency around 20 kHz to enable monitoring of relative table motion to 100 Hz
Notes: Optocad Bild in Illustrator ins Bild einarbeiten

## Slide 9: Optical layout
West table
Central table

## Slide 10: Sketch of the optical layout
Notes: Optocad Bild in Illustrator ins Bild einarbeiten

## Slide 11: Iodine stabilized Nd:YAG laser
According to the arm length mismatch
(~23 m) a highly stabilized laser is necessary.
120 Hz/sqrt(Hz) @ 10 mHz for light of 1064 nm
AEI measurements by Michael Tröbs

## Slide 12: Sketch of the optical layout
Notes: Optocad Bild in Illustrator ins Bild einarbeiten

## Slide 13
4 Interferometers
Power monitor
Monitoring of
  power fluctuations
  beam pointing
Reference IFO
Diagnostic IFO
West IFO
South IFO
Central table
West table
South table
Reference
beam
[m]
[m]
Measurement
beam
Acts as reference point
Outputs of other IFO are sub-
tracted from output of this IFO
is used to measure the
sensitivity of the SPI
monitors the motions of the
west table relative to the
central table
monitors the motions of the
south table relative to the
central table

## Slide 14: Hydroxide-catalysis bonding
Thermal drifts requires components to be monolithically bonded to plate with low CTE (Clearceram-Z HS, CTE=0.4*10-7/K)

## Slide 15: Production accruracy of the template
Δ = Difference between nominal value and
measured value
Δxmax = +0.260 mm
Δymax = +0.282 mm
Δxmin = +0.002 mm
Δymin = -0.002 mm
Position of each ball
Δxmean = (+0.095 ± 0.059) mm
Δymean = ( -0.087 ± 0.100) mm

## Slide 16: Production accruracy of the template
Results in a different
corner point
Δx = (0.041 ± 0.047) mm
Δy = (0.252 ± 0.082) mm
Δα = (0.051 ± 0.078) °
Note
a 1 µm larger ball in diameter changes the angle by 0.004 °
Δ = Difference between nominal value and
measured value

## Slide 17: Bonding accuracy

## Slide 18: Current status of the SPI

## Slide 19: TRA-BOND 2101 sold as “Vacseal“
TRA-BOND 2101
Empty vacuum chamber
TRA-BOND 2101 is not vacuum compatible

## Slide 20: MASTERBOND EP30-2 covers TRA-BOND 2101
Chart: Empty chamber minus MASTERBOND EP30-2 covers TraBond2101
Chart: Mass

## Slide 21: TRA-BOND 2151 covers TRA-BOND 2101
Chart: Mass
Chart: Empty chamber minus TraBond2151 covers TraBond2101
Two orders of magnitude less of long hydrocarbon chains
I will cover the TRA-BOND 2101
spots with TRA-BOND 2151

## Slide 22: Pointing noise of fibre injectors
Central table
South table
West table
Notes: Zeigen, wo was along the table connecting line liegt

## Slide 23: Pointing noise of fibre injectors
In simulation
reference beam at fix position
measurement beam rotated
To reach the design goal of
10 nrad, the
beam pointing
noise has to be
below 0.62 nrad

## Slide 24: Pointing noise of fibre injectors
Quadrant photodiodes
will be placed about
60 cm from the
Clearceram baseplate
Central table
South table
West table

## Slide 25: Phase determination
Phase is extracted from heterodyne
signal by use of an hardware phase
meter based on FPGA chips
Preamplifier and A/D conversion
Single bin discrete Fourier transform
  Fourier transform at only one frequency
   complex amplitude of PD signal at fhet
Data transfer from phase meter via phase meter interface (PMI) (equipped with Ethernet port) to aLIGO CDS
Signal combination of each QPD quadrant leads to phase, DC, Differential Wavefront Sensing (DWS) and contrast information

## Slide 26: Please, remember ...
... don‘t buy „Vacseal“ from Vacom. It is TRA-BOND 2101 which is not vacuum compatible.

## Slide 27
Thank you for your attention.
http://10m-prototype.aei.uni-hannover.de