## Slide 1
Christoph Affeldt
&
Stefan Hild
29.08.2011
GEO-HF EOMs

## Slide 2
Intro
Basics
Options
- 6x6x35mm³
- 8x8x10mm³
- ???
Discussion & decision!
Outline

## Slide 4
MID (15MHz, 4.25Vrms)
SR (21.7MHz, 2.73Vrms) (& power stab.)
MC2 (13MHz, 2Vrms)
PR (37MHz, 3.54Vrms)
MC1(??)
Notes: Seitenbänder…

## Slide 5: Thermal lens
Mounting Unit (MU)

## Slide 6: Thermal lens
Thermal lens & crystal destruction
We need other EOM crystals!
Notes: LIGO investigations have shown…

## Slide 7
Two options!
Two RTP (RbTiOPO4) options:
6x6x35mm³
8x8x10mm³
instead of 8x8x35mm³ LiNbO3.
Why RTP? See:
Notes: Rubidium Titanyle Phosphate r33 – electro optical coeff. nz – refractive index

## Slide 8
Decision!
Oder doch lieber Tor zwei?

## Slide 9
Basics
Two RTP (RbTiOPO4) options:
6x6x35mm³
8x8x10mm³
instead of 8x8x35mm³ LiNbO3.
Notes: Rubidium Titanyle Phosphate r33 – electro optical coeff. nz – refractive index

## Slide 10
Basics
6x6x35mm³ →
8x8x10mm³ →

## Slide 11
Option 1: 6x6x35mm³
clipping @ first Faraday: 6.4ω
With 6x6mm² aperture:
Clipping @ first EOM: 5.9ω
And:
Expensive
Never manufactured

## Slide 12
Option 2: 8x8x10mm³
Optimization of the matching circuit
Application of higher modulation voltage to the EOM crystal
Application of more than one RF modulation frequency per EOM crystal
Use of more than 2 EOM crystals on MU3

## Slide 13
MID modulation path?!

## Slide 14
Higher modulation voltage?
Yes, at least partly!
Factor 2 is easy! (just decreased MID sidebands after MSR change. Can go back from 4.25Vrms (0.4W) to 10Vrms (2W))
Commercial low noise RF amplifiers available! (e.g. “QBS-544”, but: 8W necessary, >1000€)
Notes: Infos!!

## Slide 15
Optimization of the matching circuit
Application of higher modulation voltage to the EOM crystal
Application of more than one RF modulation frequency per EOM crystal
(see also G1000297-v1.pdf K. Izumi et al. “Development of Tripple Resonant EOM for Advanced Detectors” )
Use of more than 2 EOM crystals on MU3
Option 2: 8x8x10mm³

## Slide 16
More crystals
Recovering modulation strength by using more crystals. (At least 3 necessary for MU3)
→ More potential parasitic cavities.

## Slide 17
Proposal for 8x8x10mm³
Using three 8x8x10mm³ RTP crystals in MU3
One for SR and two for MID.
Existing electronics should be capable of recovering the remaining factor of 2.24 for MID modulation.

## Slide 18
Mixing? (MU3 case – crazy?!)
One 8x8x10mm³ for SR SBs.
One 6x6x35mm³ for MID SBs
→6.16ω

## Slide 19
Let´s make a choice!
8x8x10mm³ (x2)
+ same aperture
+ Cheaper
(7000GBP)
− lower modulation (1:4.5)
− more parasitic cavities
6x6x35mm³
+ same modulation
− smaller aperture
− more expensive
(12800GBP)
− never fabricated before

## Slide 20
Other questions
What should we use in MU2?
What is going on in MU1? (Do we even need new crystals there?)

## Slide 21: Thank you for your attention!

## Slide 22
Options
8x8x10mm³
+ same aperture
+ Cheaper (7000GBP)
− lower modulation (1:4.5)
− more parasitic cavities
− needs higher modulation voltage
6x6x35mm³
+ same modulation
− smaller aperture
− never fabricated before
− more expensive
(12800GBP)
        Further questions
What should we use in MU2?
What is going on in MU1? (Do we even need new crystals there?)

## Slide 23
Costs…

## Slide 24
Decision!

## Slide 25
More than one RF modulation frequency per EOM crystal?
See K. Izumi et al. “Development of Tripple Resonant EOM for Advanced Detectors”